## Extraction Method
Extraction Experiment information
Comparative testing: Process the same Water Samples sent for Eurofin testing
Direct Extraction Details for DH5a Bacteria
Elysis |  | Steps | Conditions Used
| 1 | Add 100 volume of sample into a 1.5 mL tube | N/A
| 2 | Add 500ul Elysis buffer | N/A
| 3 | Vortex for 10seconds and spin down, before heating the tube at 80deg for 10mins | N/A
| 4 | Centrifuge solution at 15000RPM for 10mins | N/A
| 5 | Remove supernatant and add 150ul Elution buffer | N/A
Filtration method Extraction Details: MF ¦ Metal Filtration System
Elysis |  | Steps | Conditions Used
| 0 | Membrane selected | PES membrane
| 1 | Filter an appropriate volume of water sample using Metal Filtration System. | 300ml
| 2 | Place membrane in 5ml Tube with 1ml Diluent and vortex for 1min | (1 mL PBS)
| 3 | Short Spin using the centrifuge (~2000rpm) | N/A
| 4 | Transfer solution to 1.5ml microcentrifuge tube & centrifuge at 8000rpm for 5mins | N/A
| 5 | Supernatant is removed until ~100ul remains. Add 700ul Elysis buffer | N/A
| 6 | Vortex for 10seconds and heat tube at 80deg for 10mins | N/A
| 7 | Centrifuge solution at 15000RPM for 10mins | N/A
| 8 | Remove supernatant and add 150ul Elution buffer | N/A

## Reaction Set-up
Project ID: | Jaipur (EC2110-D)
Experiment Title: | 20221005 Run 1 Eurofin Samples
Experiment Aim: | To test presence of Legionella and E. coli in Eurofin Samples
Date: | 2022-10-05
Operator: | Bachana
Link:
Equipment
Description | Asset Number (Internal) | Asset Number (GIS / Serial Number) | Maintenance Due Date | Location
-20°C Freezer | 20FR-MM-1 | 9020026 | 2023-02-17 | Master Mix Room (MM)
4°C Refrigerator | 4FRG-MM-1 | 13030253 | 2023-05-04
Pipette - 1000 µL | P1000-MM-2 | KJ09437 | 2022-12-08
Pipette - 200 µL | P200-MM-2 | C006993455 | 2023-05-23
Pipette - 20 uL | P20-MM-2 | HH93920 | 2022-12-09
Pipette - 10 uL | P10-MM-2 | KJ08190 | 2022-12-08
Pipette - 2 uL | P2-MM-2 | KH43912 | 2022-12-09
Pipette - 1000 µL | P1000-TA-4 | C006992352 | 2023-05-23 | Template Addition Room (TA)
Pipette - 200 uL | P200-TA-4 | C006993052 | 2023-05-23
Pipette - 20 uL | P20-TA-4 | OZ00190 | 2023-05-23
Pipette - 10 uL | P10-TA-4 | C006992852 | 2023-05-23
Real-Time PCR System, Rotor-Gene Q 5 Plex HRM System | CFX 96 (IVD) | BR101990 | 2022-06-02 | Instrument Room
Microplate Centrifuge | SP-BIO-NPC-1 | TT-E14-1132 | NA
Reagents
Description | Vendor | Catalog Number | Lot Number | Date of Expiry /Manufacture
Leg_mipT1_F26 | IDT | N/A | 105977862 | 2022-03-10
Leg_mipT1_R23 | IDT | N/A | 105977863 | 2022-03-10
Leg_mip_P1_FAM | IDT | N/A | 105995262 | 2022-03-14
Legio_16s_F35 | IDT | N/A | 105954788 | 2022-03-03
Legio_16s_R23 | IDT | N/A | 105848683 | 2022-02-11
Legio_16s_R24 | IDT | N/A | 105848684 | 2022-02-11
Leg_16S-P2 HEX | IDT | N/A | 105976752 | 2022-03-14
Leg_wzm_F4 | IDT | N/A | 105845947 | 2022-02-10
Leg_wzm_R9 | IDT | N/A | 105845962 | 2022-02-10
Leg_wzm_P2-TX | IDT | N/A | 106374997 | 2022-06-22
2x SensiFAST Probe No ROX Mix | MERIDIAN | NA | SF593-B092150 | DOE: MAR 2023
Nuclease Free Water | QIAGEN | 163041698 | 1039498 | NA
UidA-F4 | IDT | NA | 106156521 | 2022-04-22
UidA-R2 | IDT | NA | 106156525 | 2022-04-22
UidA-P1 FAM | IDT | NA | 106292280 | 2022-05-12
Samples
Description | Vendor | Catalog Number | Lot Number | Date of Expiry /Manufacture
Legionella Pneumophila S14 100kx | ATCC | 43703 | 70036189 | 2022-01-20
Legionella Pneumophila S1e2 | ATCC | 33152DQ | 70036289 | 2022-03-08
Consumables
Description | Vendor | Catalog Number
Pipette Tips - 1000 µL | Axygen | TF-1000-L-R-S
Pipette Tips - 200 µL |  | TF-200-L-R-S
Pipette Tips - 20 µL |  | TF-20-L-R-S
Pipette Tips - 10 µL | Axygen | TF-300-L-R-S
Microtubes, 1.5 mL |  | MCT-150-C
Microtubes, 2.0 mL |  | MCT-200-C
Hard-Shell 96-Well PCR Plates, Low Profile, Thin Wall, Skirted, White/Clear | Bio-Rad | HSP9601
Microseal 'B' seal; Seal | Bio-Rad | MSB1001
Master Mix Room
Primer dilution
Reagents | Stock (uM) | Final (uM) | Dilution Factor | Amount of Primer | Amount of Water/TE | Final Volume (ul)
Primers | 100 | 10 | 10 | 10 | 90 | #REF!
uidA_F4R2P1 |  |  |  |  | Legionella 3plex: MipT1F26R23P1 16SF45R27P2 WZMF4R9P2
Component |  | Per Reaction (uL) | 18 |  | Component |  | Per Reaction (uL) | 18
2x SensiFAST Probe No ROX Mix |  | 10 | 180 |  | 2x SensiFAST Probe No ROX Mix |  | 10 | 180
UidA_F4 (10uM) |  | 0.5 | 9 |  | Leg_mipT1_F26 (10uM) |  | 0.5 | 9
UidA_R2 (10uM) |  | 0.5 | 9 |  | Leg_mipT1_R23 (10uM) |  | 0.5 | 9
uidA-P1 FAM (10uM) |  | 0.5 | 9 |  | Leg_mipT1_P1_FAM (10uM) |  | 0.5 | 9
H2O |  | 3.5 | 63 |  | Leg_16s (B2)_F45 (10uM) |  | 0.5 | 9
Total |  | 15 | 270 |  | Leg_16s (B2)_R27 (10uM) |  | 0.5 | 9
[5 empty] | Leg_16s (B2)_P2_HEX (10uM) |  | 0.5 | 9
[5 empty] | Leg_wzm_F4 (10uM) |  | 0.625 | 11.25
[5 empty] | Leg_wzm_R9 (10uM) |  | 0.625 | 11.25
[5 empty] | Leg_wzm_P2_TEX (10uM) |  | 0.5 | 9
[5 empty] | H2O |  | 0.25 | 4.5
[5 empty] | Total |  | 15 | 270
Template Addition Room
Add 5uL extracted samples
- Add 15µL MM + 5uL template into respective wells in accordance to the layout below
Cycling Parameters
Step | Stage | Number of Cycles | Temperature (°C) | Time (Min : Sec) | Acquiring
Hold | PCR Initial Heat Inactivation | 1 | 95 | 2:00 | -
Cycling | Denaturation | 45 | 95 | 5 sec | -
| Annealing + Extension |  | 60 | 30 sec | Acquiring to all channels

## Plate Layout
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A |  | UidA A1 MF PES 300ml NC | UidA A1 MF PES 300ml NC |  | 3plx A1 MF PES 300ml NC | 3plx A1 MF PES 300ml NC
B | [9 empty] | UidA NTC | UidA NTC
C |  | UidA A1 MF PES 300ml L8F Tapwater | UidA A1 MF PES 300ml L8F Tapwater |  | 3plx A1 MF PES 300ml L8F Tapwater | 3plx A1 MF PES 300ml L8F Tapwater |  |  |  | 3Plex A1 MF PES NTC | 3Plex A1 MF PES NTC
D |  | UidA A1 MF PES 300ml L7F Shower | UidA A1 MF PES 300ml L7F Shower |  | 3plx A1 MF PES 300ml L7F Shower | 3plx A1 MF PES 300ml L7F Shower
E |  | UidA A1 MF PES 300ml Matrix | UidA A1 MF PES 300ml Matrix |  | 3plx A1 MF PES 300ml Matrix | 3plx A1 MF PES 300ml Matrix
F |  | UidA A1 MF PES 300ml L7 Water Dispenser | UidA A1 MF PES 300ml L7 Water Dispenser |  | 3plx A1 MF PES 300ml L7 Water Dispenser | 3plx A1 MF PES 300ml L7 Water Dispenser
G |  | UidA A1 MF PES 300ml L7 Koi Pond | UidA A1 MF PES 300ml L7 Koi Pond |  | 3plx A1 MF PES 300ml L7 Koi Pond | 3plx A1 MF PES 300ml L7 Koi Pond
H |  | UidA A1 MF PES 300ml EC Spiked | UidA A1 MF PES 300ml EC Spiked |  | 3plx A1 MF PES 300ml EC Spiked | 3plx A1 MF PES 300ml EC Spiked

## Run Analysis
Project ID: | Jaipur
Experiment Title: | 20220817 Run 2 EC, leg 3plex w 700ml Ice AF crude
Experiment Aim: | To check NTC for 2/3PLEX by using different primer combination
Date: | 2022-08-17
Operator: | Bachana
Link:
Instrument Information
Description |  | Asset Number (Internal)
Real-Time PCR System, CFX96 Touch |  | CFX-INS-2
Analysis Settings
Target Channel: | FAM, HEX, TEX
Threshold (CFX): | 50
Results and Run Analysis
Raw Amplification Plot (Green (FAM)) |  |  |  |  | Raw Amplification Plot (Yellow (HEX))
Well | Sample | Ct in FAM channel | Ct in HEX channel | Ct in TEX channel | Raw Amplification Plot (Red (TEX))
A02 | UidA MF 300mlNC | N/A | N/A | N/A
A03 | UidA MF 300mlNC | N/A | N/A | N/A
A05 | 3plx MF300ml NC | N/A | N/A | N/A
A06 | 3plx MF300ml NC | N/A | N/A | N/A
B10 | UIdA NTC | N/A | N/A | N/A
B11 | UIdA NTC | N/A | N/A | N/A
C02 | UidA MF L8F Tap | N/A | N/A | N/A
C03 | UidA MF L8F Tap | N/A | N/A | N/A
C05 | 3plex MF L8FTAP | 30.36 | 28.49 | 30.47
C06 | 3plex MF L8FTAP | 30.6 | 28.39 | 30.86
C10 | 3Plex NTC | N/A | N/A | N/A
C11 | 3Plex NTC | N/A | N/A | N/A
D02 | UidA MF L7FShwr | N/A | N/A | N/A
D03 | UidA MF L7FShwr | N/A | N/A | N/A
D05 | 3plxMF L7FShowr | 28.74 | 26.5 | 29.16
D06 | 3plxMF L7FShowr | 28.58 | 26.8 | 29.04
E02 | UidA MF Matrix | N/A | N/A | N/A
E03 | UidA MF Matrix | N/A | N/A | N/A
E05 | 3plxMF Matrix | 30.08 | 28.94 | N/A
E06 | 3plxMF Matrix | 29.98 | 29.1 | N/A
F02 | UidA MF L7 Disp | N/A | N/A | N/A
F03 | UidA MF L7 Disp | N/A | N/A | N/A
F05 | 3plxMF L7Dispen | N/A | 38.44 | N/A
F06 | 3plxMF L7Dispen | N/A | 35.52 | N/A
G02 | UidA MF Koipnd | N/A | N/A | N/A
G03 | UidA MF Koipnd | N/A | N/A | N/A
G05 | 3plx MF KoiPnd | N/A | 25.46 | N/A
G06 | 3plx MF KoiPnd | 32.46 | 24.47 | 36.52
H02 | UidAMF SpikeEC | 28.92 | N/A | N/A
H03 | UidAMF SpikeEC | 28.71 | N/A | N/A
H05 | 3plxMF SpikedEC | N/A | 38.52 | N/A
H06 | 3plxMF SpikedEC | N/A | 40.12 | N/A